 Ventilator Troubleshooting Alarms
Skills Checklist
	Child’s Name: _______________________________ 

	Person Trained: _____________________________ Position: _____________________ Instructor: _____________________



	
	Demo
	Return Demonstration

	
	Date
	Date
	Date
	Date
	Date
	Date
	Date

	A.	States name and purpose of procedure.
	
	
	
	
	
	
	

	B.	Steps:
	1.	Identifies which alarm is sounding
	
	
	
	
	
	
	

		2.	Checks child first, if LOW PRESSURE alarm sounds
	
	
	
	
	
	
	

	1. Removes child from ventilator and gives breaths with resuscitator bag
	
	
	
	
	
	
	

		4.	Checks for leaks; if child is fine:
	
	
	
	
	
	
	

	a. Child disconnected
	
	
	
	
	
	
	

	b. Disconnected tubing
	
	
	
	
	
	
	

	c. Kinked tubing
	
	
	
	
	
	
	

	d. Punctured tubing
	
	
	
	
	
	
	

	e. Water in exhalation valve
	
	
	
	
	
	
	

	f. Hole in exhalation valve
	
	
	
	
	
	
	

	g. Loose fitting heater humidifier
	
	
	
	
	
	
	

	h. Accidental change in ventilator settings
	
	
	
	
	
	
	

		5.	Tests system after leak is found
	
	
	
	
	
	
	

		6.	Places child back on ventilator
	
	
	
	
	
	
	

		7.	Checks child first if a HIGH PRESSURE alarm sounds
	
	
	
	
	
	
	

		8.	Checks activity of child:
	
	
	
	
	
	
	

	a. Coughing
	
	
	
	
	
	
	

	b. Sneezing
	
	
	
	
	
	
	

	c. Talking
	
	
	
	
	
	
	

	d. Laughing
	
	
	
	
	
	
	

	e. Crying
	
	
	
	
	
	
	

	f. Hiccups
	
	
	
	
	
	
	

	g. Body position
	
	
	
	
	
	
	

	h. Breath holding
	
	
	
	
	
	
	

		9.	Suctions, if needed
	
	
	
	
	
	
	

		10.	Changes trach tube, if needed
	
	
	
	
	
	
	

	11.   Removes child from ventilator and gives breaths with
        resuscitator bag
	
	
	
	
	
	
	

		12.	Checks ventilator for obstructions, if child is okay:
	
	
	
	
	
	
	

	a. Kinks in tubing
	
	
	
	
	
	
	

	b. Water in tubing
	
	
	
	
	
	
	

	c. Blocked exhalation valve
	
	
	
	
	
	
	

	d. Accidental change in ventilator settings
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		13.	Places child back on ventilator once problem is solved
	
	
	
	
	
	
	

		14.	Checks below if POWER SOURCE alarm:
	
	
	
	
	
	
	

	a. AC power
	
	
	
	
	
	
	

	b. Internal battery
	
	
	
	
	
	
	

	c. External battery
	
	
	
	
	
	
	

	15. Removes child from ventilator if all 3 systems fail and gives
breaths with resuscitator bag
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